US-191 "IMAGO" - Direction de l'US191 et la composante océanographique de Bretagne

User Manual

for

TSGQC

ThermoSalinoGraph Quality Control

Version 1.5

| Rédacteur |  | Approbateur |  |
| --- | --- | --- | --- |
| Nom : | Yves Gouriou Jacques Grelet Gaël Alory | Nom : |  |
| Fonction : |  | Fonction : |  |

# Revision History

| Dates | Versions | Chapitres concernés | Commentaires et modifications |
| --- | --- | --- | --- |
| August 2010 | 1.5 | Creation | English version |

# Contents

1.	Revision History	2

2.	Contents	3

3.	Software goals	4

4.	A Quick Start	5

5.	Functionalities	6

A.	Starting TSGQC	6

B.	Main functions	7

1.	Toolbar	7

2.	Menus	8

C.	Opening a File	11

1.	Opening a « Merchant Vessel » TSG file (SODA format)	13

2.	Opening Sample Files (*.btl, *.spl, *argo)	15

D.	Saving your work	16

1.	netCDF format	16

2.	ASCII format	17

E.	Using TSGQC	19

1.	Parameters visualization	19

2.	Plotting a climatology	21

3.	Plotting a Map	22

4.	Zoom and Pan functions	23

5.	Selecting time limits	24

6.	Printing figures	25

F.	Attributing Quality codes to the TSG measurements	26

G.	Calibration	28

H.	Interpolating latitudes and longitudes	29

I.	Correcting the data	30

1.	Description	31

2.	Functionalities	32

3.	Algorithms	35

J.	Data processing report	37

K.	Quit TSGQC	39

6.	Data Format	40

A.	Name of the variables	40

1.	Main variables used for TSG data	40

2.	Discrete sample	41

B.	Input files	42

C.	Output files	42

D.	ARGO Format (*.arg)	42

E.	ASTROLABE Format (*.ast)	43

F.	ASCII Format (*.btl, *.spl, *.tsgqc)	43

G.	netCDF format (*.nc)	44

H.	SODA format (*.lbv)	44

I.	NUKA ARTICA format (*.transmit)	44

J.	Climatology file format	45

K.	Glossaire	47

# Software goals

TSGQC (ThermoSalinoGraph Quality Control) software was build to validate, calibrate and correct salinity and temperature measurements carried out with thermosalinograph on board merchant ship and research vessels:

- Validate:	Quality Control code assigned to every measurement.

- Calibrate:	Apply a linear shift to sensor measurement

- Correct:	Use of discrete sample as water sample, ARGO, CTD, XCTD measurements to correct the TSG time series.

# A Quick Start

The main steps to quality control a TSG file are:

- Starting TSGQC	p. 6

- Open a TSG file. Several formats are available	p. 12

- Visual screening of the data

  - Compare  the measurements to a climatology	p. 21

  - Plot the ship track	p. 23

- Check that every measurement has a position	p. Error: Reference source not found

- Assign a quality code to every measurement	p. Error: Reference source not found

- Saving your work	p. 17

  - netCDF format	p. 18

  - ASCII format	p. 18

- Open sample files	p. 16

- Adjust the TSG data to the sample measurements	p. Error: Reference source not found

- Printing figures	p. Error: Reference source not found

Secondary TSGQC functions:

- Using automatic tests	p. 8

- Measurements calibration	p. Error: Reference source not found

- Data treatment report	p. 38

# Functionalities

## Starting TSGQC

Launch TSGQC in the Matlab 'Command Window'.

The data directory can be chosen in the list-selection dialog box.

Type the name of the program in the command window:

>> tsgqc

The software functionalities can be used once a TSG file has been read (p. Error: Reference source not found)

## Main functions

### Toolbar

Opening files	p. 12

Writing files in netCDF	p. 17

Printing figures	p. Error: Reference source not found

Zoom in	p. 24

Zoom out	p. 24

Enable and disable pan mode on a plot	p. 24

Manual Quality Control module	p. Error: Reference source not found

Select date limits	p. 25

Vessel track figure	p. Error: Reference source not found

Plotting climatology	p. Error: Reference source not found

Calibration module	p. Error: Reference source not found

Interpolate position	p. Error: Reference source not found

Adjusting TSG to sample measurements	p. Error: Reference source not found

Meta-Data form	p. 17

Report files	p. 38

### Menus

FILE menu

Open	Open a file (TSG, samples, etc,)	p. 12

Save	Saving the data in a NetCDF format	p. 17

Export	Saving the data in an ASCII format	p. 18

    - TSG measurements

    - Samples (Water Samples, CTD, ARGO, …)

Quit	Quit TSGQC	p. 40

Edit menu

This menu is not implemented.

Climatology menu

This menu allows the user to choose the climatology period that will be superimposed on the time series measurements:

1. Yearly

2. Seasonally

3. Monthly

The climatology and its depth can be selected through the Option - Preferences menu (see next paragraph)

Option menu

The Option - references menu can be used to set some preferences:

- Climatology and its depth	p. 21

- Calibration coefficient types

- Figures

- Automatic tests limits that will be used with the QCauto menu (next paragraph)

QCauto menu

QCauto allows the user to apply automatic tests to the TSG time series. The tests use limits defined in the “Option - Preferences” form (see previous paragraph).

A Quality Code “BAD” is applied to measurements outside the limits defined for the test.

Remarks

- There is no Undo function.

- Quality code “HARBOUR” is applied to measurements carried out when the speed of the vessel is below the “Ship Speed Min” value.

- Be careful with automatic tests. For example when a vessel is stopped, measurements can be BAD simply because there is no more water pressure in the TSG. In that case the quality code should be set to “BAD” and not to “HARBOUR”.

Help menu

Online help is not yet implemented

## Opening a File

Several file types can be read in TSGQC:

*.arg	ARGO data file created daily by G. Reverdin team (UMR LOCEAN) close to the ship track

*.ast	TSG file from the ASTROLABE

*.btl	Water Sample file (ASCII format)

*.lbv	TSG file (SODA format) from merchant ships (SOERE SSS)

*.nc	TSG file (netCDF format)

*.spl	‘Sample’ files (XBT, CTD, etc. - ASCII format)

*.transmit	TSG file from the merchant ship NUKA ARTICA

*.tsgqc	TSG file (ASCII format)

Format description can be found page:

*.arg	p. 43

*.ast	p. 44

*.btl	p. 44

*.lbv	p. 45

*.nc	p. 45

*.spl	p. 44

*.transmit	p. 46

*.tsgqc	p. 44

Remark

A test on double records is performed for:

- TSG measurements: If consecutive dates are detected the following message box is displayed. The message gives the number of double records et display the 10 first identical dates. Double records are deleted.

- Sample measurements: Double measurements are detected based on date and hour comparison. No message box is displayed. Double records are deleted.

Sample file (*.btl, *.spl, *.arg) cannot be read if no TSG file was read. The following message box is displayed:

### Opening a « Merchant Vessel » TSG file (SODA format)

When opening a Merchant Vessel file in the SODA format, the following window is displayed:

This window gives a list of the variables present in the header line. TSGQC selects the variables essential to its running. The user can modify these choices. For example it is better to choose a GPS date and hour than a PC time.

sbe21_raw variable stores frequency measurements made by the Sea-Bird TSG. This information is recorder in netCDF file.

sbe21_ad1 variable stores flow measurements. Data are copied in the netCDF FLOW variable.

Remark

The vessel speed velocity stores in SODA file (variable SOG) is truncated. It is better to uncheck this variable and let the software compute the speed velocity from measurements position.

### Opening Sample Files (*.btl, *.spl, *argo)

Samples (variable with extension _EXT in the netCDF format) can be used to correct, adjust the TSG measurements.

It is possible to open and concatenate several sample files. A message box asks the user if he wants to replace or concatenate the sample file already loaded in memory, if it exists.

Remarque

- If two or more files are concatenated TSGQC test for double records. The test is made on the date (year, month, day, hour, minute, second). Double records are deleted. No message box indicating that double records were detected is displayed.

- Sample files (*.btl, *.spl, *.arg) cannot be read if there is no TSG file in memory. In that case the following message box is displayed:

## Saving your work

### netCDF format

The default recording format is netCDF (see p. 45). Every information (raw data, calibrated data, adjusted data, meta-data) is stored in the netCDF file.

Metadata are either read in the original file or the user fills the following input form.

This form is automatically displayed when recording a netCDF file. The form is also available through the shortcut

PLATFORM_NAME is the only information necessary to save a file. In general use the ship name in upper-case letters.

For a given PLATFORM_NAME some metadata are saved in an ASCII file. As soon as the user enters a PLATFORM_NAME, TSGQC scan the ASCII file to get all the information about this PLATFORM_NAME.

Jacques peut-on copier le fichier tsg_platform sur un autre ordinateur pour ne pas avoir à entrer à nouveau toutes les informations suite à un crash de disque par exemple.

It is recommended to regularly save your work.

### ASCII format

This format can be used for TSG and Sample measurements. To save data in the ASCII format use the File - Export menu.

The user can export :

- Data samples (extension name: .spl): A .spl file contains all data samples loaded in TSGQC. For the ARGO file the mean time difference between the TSG and ARGO measurements is added to the TSG date (column 17 of the *.arg file see p. 43). It is this date which is recorded in the netCDF and ASCII file. TSGQC proposes to export the following variables:

- TSG data (file with extension .tsgqc). TSGQC proposes to export the following variables:

A description of the ASCII format is given p. 44.

## Using TSGQC

### Parameters visualization

Once a TSG file has been loaded TSGQC displays 3 figures. The default variables plotted in the figures are:

- SSPS: 	The salinity in the upper plot. This plot is used to interactively attributes Quality codes to the data.

- SSJT: 	Jacket TSG temperature in the middle figure.

- SPDC: 	The vessel speed velocity in the bottom figure.

When the user moves the pointer on the TSG curve, information on the date, position, salinity and temperature are displayed in the upper banner

The user can choose the parameters that will be plotted in the three figures. For this, use the three list-selection dialog boxes in the frame box ‘TSG Parameter’ on the left of the screen.

Parameters:

SSPS	Computed salinity using NDC and SSJT.

SSJT	Jacket temperature measured by the TSG.

SSTP	Temperature given by a sensor different from the TSG (SBE 38 for example)

SPDC	Ship speed velocity.

CNDC	Conductivity measured by the TSG

SSPS_STD	For average or median TSG records this variable gives the standard deviation of the mean value.

FLOW	Sea water flow in the TSG.

PRES	Water pressure in the TSG

LATX	Latitude en decimal degree

LONX	Longitude en decimal degree

Remark

If the ship speed velocity is not available in the TSG file, it is computed by using TSG measurement positions.

### Plotting a climatology

The user can choose climatology in the menu Option - Preferences. The only available climatology are Levitus WOA01 and WOA05. Climatology depth can also be modified in this menu. The user can choose between 0 and 10 m.

The mean climatology value (black line) and minus/plus 3 times the standard deviation is superimposed on the TSG measurement curve (salinity or temperature).

By default, the annual climatology is displayed. The user can display seasonal or monthly climatologies using the 'Climatology' menu (see p. 8).

The format of the climatology files is described p. 46

### Plotting a Map

Click the following shortcut         to get the ship track plotted on a Mercator projection map.

The pointer position on the TSG time series (upper plot) is reported as a red circle on the cruise track. This is useful to visualize the position of the TSG measurement.

The plot is made taking into account the temporal limits of the upper plot: if the user zooms in the TSG time series, the map is also zoomed in. The map limits are updated automatically.

The quality code colour assigned to each TSG measurements are used to visualise the cruise track.

Several map resolutions are available. Use the menu « Option - Preferences ».

The map is plotted in a Matlab figure window. The figure menus can be used to print or export the map in a graphic file.

The map can be displayed is several TSGQC modules..

### Zoom and Pan functions

Zoom tools : The user can increase or increase part of the figure. Zooming is useful to see greater detail in a small area. When in zoom in mode, the user can draw a rectangle with the pointer around the time series to magnify.

The 3 curves undergo the same change.

You can move your view of a graph up and down as well as left and right with the Pan tool. Panning is useful when you have zoomed in on a graph and want to translate the plot to view different portions.

Les 3 plot undergo the same displacement.

In Zoom and Pan modes, if you want to reset the graph to its original view, right-click to display the context menu and select Reset to Original View.

### Selecting time limits

The user can select part of the time series by setting time limits in the left panel ‘Date Limits’. The selection can be made:

- Manually by using the input dialog box of the ‘Date Limits’ panel.

- With the mouse, by selecting a portion of the time series. The time limits of the rectangle defined by the pointer are entered automatically in the input dialog box.

This tool can be used in different modules of TSGQC (position interpolation, correction of the data, etc.)

### Printing figures

This is a simplified tool to print TSGQC plots. Whatever the TSGQC module you are working on, you can print the plots displayed on the screen by clicking on the icon       .The plots will be drawn inside a Matlab Figure Window.

Then you can use the Matlab menus to print, save or export the plots.

Note that the cruise track is plotted directly in Matlab figure window.

## Attributing Quality codes to the TSG measurements

The user can assign a quality code only to the measurements of the time series displayed in the upper plot. Quality code can be attributed to the 3 variables SSPS, SSJT, and SSTP. You have to display the variable that you want to assign the quality code (see p. 20)

Enabling the validation module allows displaying quality codes on the left of the screen, inside the “Validation Codes” panel. The number to the left of the radio buttons give the number of measurements assign to peculiar code (7677 have a quality code GOOD in that example).

When the user moves the cross-shape pointer on the time series, he plots a rectangle which sets the limits of the data measurements whose quality codes will be modified.  The code is chosen with 2 methods:

- Before the selection, activate the radio button (click the mouse button on the object) in the “Validation Codes” panel.

- After the selection, right-click to display the following context menu:

Selected data measurements get the colour of the chosen code. The HARBOUR code is designed to mark good data taken when the ship is in a port. If the data are not of a good quality choose another code.

It is of good practice to zoom in the time series, to work with precision.

You can plot a vertical line that span the three plots to spot peculiar features:

- First left-click on the upper plot.

- Second press the Ctrl Key. appuyer sur la touche Ctrl : The vertical line is plotted. It disappear when the user release the key. Then move the pointer to another position and press the Ctrl Key. It is no more necessary to left-click.

Important

- There is a shortcut to Pan the figure while you are attributing a quality code to the measurements. The Pan function is available when you move the pointer in the bottom part of the upper plot. The pointer shape changes from a cross shape to a hand shape, indicating that the Pan function is activated.

- If the map is displayed when you execute a Pan operation, the map is redrawn and takes into account the new limits.

- When all the quality codes have been attributed it is important to check that the number of No Control codes is equal to 0.

## Calibration

You can apply a linear shift (slope and offset) to the conductivity (CNDC) and jacket temperature (SSJT) time series or to part of the time series.

Enter the slope and offset in the input dialog box in the Calibration panel for:

- The conductivity

- The jacket temperature (temperature measured by the TSG)

- Precise Temperature if a sensor exists.

Once these data entered click on the ‘Calibrate’ button.

Non-calibrated data are plotted in black, calibrated data in red.

Use the ‘Cancel Calibration’ button to cancel the corrections.

Remark

That function is not frequently used. It is better to use Samples to correct/Adjust the data.

## Interpolating latitudes and longitudes

Interruption of GPS acquisition can lead to get salinity and temperature measurements without any position (but referenced in time from the PC for example).

A latitude and longitude can be assigned to this measurement using a linear interpolation between known positions. This is a crude method which supposes that the ship has a constant velocity.

The method is satisfactory in the majority of cases. The interpolation figure window is the following:

The upper plot displays the salinity time series. Blue points indicate measurements without position. The middle plot display the latitude of the measurements and the bottom plot the longitudes.

Once the user have clicked on the Linear Interpolation button, interpolated latitudes and longitudes are drawn in red (plots 2 and 3). To check the validity of the interpolation it is better to zoom in. The quality code of the latitude and longitude (POSITION_QC) is set to 8 (code: Interpolated value).

You can make the interpolation only on some section of the ship track. In that case use the Date Limits tools to select the temporal limits (see p 25.

## Correcting the data

TSG data can be corrected by comparison with discrete samples (Water Sample, CTD, XCTD, ARGO...).

The parameter (SSPS, SSTP...) that will be corrected should have been first displayed in the upper plot of the first screen of TSGQC (see p. 20 ).

Data corrected are stored in variables with the extension _ADJUSTED: SSPS_ADJUSTED, SSTP_ADJUSTED, .... The raw data are not modified.

If corrections are not applied to the entire time series but only to some part of the measurements, the _ADJUSTED variables only contain corrected data. The other elements of the _ADJUSTED variables are set to NaN (Not a Number).

An error is assigned to every corrected measure (see the algorithms p. 36). Error values are stored in variables with the extension _ADJUSTED_ERROR.

The quality code of the adjusted data, _ADJUSTED_QC, is identical to the quality code assigned to the raw variable (_QC).

N’aurions nous pas du positionner le code qualité à 5 VALUE_CHANGED pour les mesures corrigées ?

Here is the Validation module window:

### Description

Upper plot

Differences between discrete samples and TSG data are drawn in the upper plot.

Square markers are used to plot Water samples collected on board merchant ship (files with extension *.btl). Circle markers are used for the other samples (ARGO, CTD ... - files with extension *.spl, *.arg).

The blue curve is the correction that will be applied to the TSG data. The correction can be made in several parts and it is not necessary applied on the entire time series.

Quality codes can be attributed to samples with the same method than for the TSG time series (see p. 27).

Remark

If for some samples there is no TSG measurements close in time or if the TSG measurements have quality code sets to Probably Bad or Bad the differences are not drawn in the upper plot.

Middle plot

The raw TSG time series colored with its quality codes is drawn in this figure.

Bottom plot

In this figure the following elements are drawn :

- In black, the raw TSG time series

- In red, the adjusted TSG time series.

  - Either after being calibrated.

  - Or after comparison with discrete samples. Adjusted data from comparison with discrete samples replace calibrated data.

- In green, the error on the adjusted TSG time series.

### Functionalities

Quality codes of discrete samples

Quality codes can be attributed to samples with the same method than for the TSG time series (see p. 27).

Only discrete samples with quality codes set to NO_CONTROL, GOOD, and PROBABLY_GOOD are used to compute the correction applied to the TSG time series.

Choice of date limits

The correction can be applied on the entire time series or to parts of it. This is done using the time selection tool (see p 25).

Correction depending on the quality code

Correction is applied to the raw TSG time series depending on the quality codes. To select the quality codes that will determine which part of the time series will be adjusted, click the check boxes in the panel figure Correction applied to....

In the following example the correction will be applied only to TSG measurements with a Good and Probably Good quality code.

Correction Methods

Four different methods of correction can be applied to the TSG time series (or part of). The user can apply different methods of corrections on different parts of the time series.

You can cancel all the corrections or part of them.

Running median filter.  See the algorithm p. 36

Detail of the algorithm is described in the following document:

G. Reverdin, F. Gaillard, S. Contardo, D. Mathias, Y. Gouriou, D. Dagorne, Réseau bleu Coriolis : Qualification des mesures navires/ Salinité de surface, Coriolis, pp 60, septembre 2006.

Linear adjustment, to correct instrumental drift mainly due to biofouling of the sensor.

See the algorithm p. 37

Gradient adjustment, to smoothly correct TSG time series in regions of strong salinity/temperature gradient where the TSG/sample difference shifts sharply across the gradient

See the algorithm p. Error: Reference source not found

Bias adjustment, to apply a constant correction estimated from a climatology or other trips, when no sample is available.

Enter the bias in the following input dialog box:

The push-button ‘Cancel within Date Limits’ can be used to cancel the correction applied on section of the TSG time series. The section is defined by its time limits (panel ‘Date Limits’).

The time-limits selection is made using the Date limits tool (see p 25).

The push-button ‘Cancel all’ cancel all the correction applied to the time series.

### Algorithms

#### Median filter correction

Step 1: Difference between TSG and discrete sample data

- Data are collocated in time. TSGQC selects the TSG measurements the closest of the sample. Function: diffTsgSample.m.

- The sample data are not compared to the collocated TSG measurements but to a mean of TSG measurement around the collocated TSG data. Function: tsg_average.m.

  - The TSG average is computed with measurements with a quality code: NO_CONTROL, GOOD, and PROBABLY_GOOD.

  - The window used to calculate the average is given by the variable tsg.cst.TSG_DT_SMOOTH. This variable is defined in the tsg_initialisation.m function. The window is equal to 1 hour.

  - A rejection test of the TSG data, based on the standard deviation, is performed to exclude TSG data with too much dispersion. TSG data which deviate more than STD_MAX from the average are excluded. This value is defined in the tsg_initialisation.m function. It is equal to 0 for the salinity and 1°C for the temperature.

- TSG-sample difference is computed as soon as the 2 measurements are not more remote from TSG_WS_TIMEDIFF. This value is defined in the tsg_initialisation.m function. It is equal to 5 minutes.

Surface salinity is a quantity that can vary quickly; the comparison of TSG data with water samples, whose dating is not always very precise, is complicated by the presence of salinity peaks, even if they are real. That is why TSG data are smoothed over 1 hour time period. We thus avoid comparing water samples to outlying salinity measurements (non representative) and then to produce a bad estimation of the correction to be applied to the TSG salinities.

Step 2: Correction

- We only take into account samples with a quality code equal to: NO_CONTROL, GOOD, and PROBABLY_GOOD.

- For each point of comparison TSGQC computes the mean value and the standard deviation of the differences TSG minus Samples in a COR_TIME_WINDOWS window. This value is defined in the tsg_initialisation.m function. It is equal to 10 days. Samples deviating more than 3 times the standard deviation from the average are eliminated. Their quality code is set to BAD.

- For each comparison point TSGQC computes the median value of the differences TSG-Samples in the window COR_TIME_WINDOWS.

- The median values of the differences are then linearly interpolated at the date of the TSG measurements and the correction is added to the raw TSG data.

Step 3: Error calculation

An estimated correction error is then computed for each TSG measurement. The error is a function of the number of point of comparison (= N) used to compute the correction value:

• if N < 4 : error = 1.

• if N ≥ 4: error = standard deviation (TSG - sample) / √(N-1) )

In no case the error cannot be less than 0.01. If errors lower than 0.01 are computed they are not kept.

#### Linear adjustment

Step 1: Difference between TSG and discrete sample data

Cf. paragraph a) Median filter correction p. 36

Step 2 : Correction

TSGQC uses the Matlab function:        [p, S, mu] = polyfit(X, Y, 1);

The linear adjustment is made with at minimum 3 samples.

If there are only 2 samples, TSGQC computes a mean difference.

Step 3: Error calculation

The error of the linear adjustment is computed using the Matlab function:

[y,error]= polyval(p,x,S,mu)

In the case of there are only 2 samples, the error is equal to the mean difference divided by 2.

In no case the error cannot be less than 0.01. If errors lower than 0.01 are computed they are not kept.

#### Gradient adjustment

Step 1: Difference between TSG and discrete sample data

Cf. paragraph a) Median filter correction p. 36

Step 2: Correction

Between two consecutive samples in the correction window where TSG-sample differences are D1 and D2, the correction value D varies from D1 to D2 as a linear function of the salinity (or temperature) variations of the TSG.

Step 3: Error calculation

The default error value is half the correction value. If the error computed this way is smaller than the sensor accuracy (about 0.01), then the later is used as the error value.

## Data processing report

Click this icon to get a simplified report on the processing of the TSG.

The report file has the following format:

TSGQC REPORT

03-Aug-2010 10:06:08

TSG file : D:\svn\tsg-qc\tsg_data\touc0702.nc

No water sample file used during this session

No external sample file used during this session

14684 total number of records

0 records have interpolated position

0 records have been deleted because they have no date

0 records deleted because their date are not increasing

0 records deleted because of velocity > 50 knots

****************** CNDC PARAMETER **************

Time series not calibrated

Number of measurements : 14684

****************** SSPS PARAMETER **************

Time series not calibrated

Number of measurements : 14684

43 -  0.29 % NO_CONTROL code

11445 - 77.94 % GOOD code

355 -  2.42 % PROBABLY_GOOD code

433 -  2.95 % PROBABLY_BAD code

1633 - 11.12 % BAD code

0 -  0.00 % VALUE_CHANGED code

754 -  5.13 % HARBOUR code

0 -  0.00 % NOT_USED code

0 -  0.00 % INTERPOLATED_VALUE code

21 -  0.14 % MISSING_VALUE code

6504 - 44.29 % records have been corrected

9 ARGO samples in the file

24   WS samples in the file

No  CTD sample

No  XBT sample

****************** SSJT PARAMETER **************

Time series not calibrated

Number of measurements : 14684

14663 - 99.86 % NO_CONTROL code

0 -  0.00 % GOOD code

0 -  0.00 % PROBABLY_GOOD code

0 -  0.00 % PROBABLY_BAD code

0 -  0.00 % BAD code

0 -  0.00 % VALUE_CHANGED code

0 -  0.00 % HARBOUR code

0 -  0.00 % NOT_USED code

0 -  0.00 % INTERPOLATED_VALUE code

21 -  0.14 % MISSING_VALUE code

No records have been adjusted

****************** SSTP PARAMETER **************

no time series

## Quit TSGQC

If the TSG measurements have been modified (quality codes, adjustment, etc.) and if the file has not been saved the following message box is displayed:

# Data Format

## Name of the variables

The variable names use in TSGQC follow the ROSCOP Parameter Code. All the variables are described in the CORTSG_format_gosud.doc document.

The same variable names are used to write and read the header of the ASCII files.

TSGQC makes a distinction between high resolution TSG data (< 5 minutes) and discrete sample data (Water samples, CTD, ARGO, XBT, XCTD, etc.).

### Main variables used for TSG data

The main variables used to describe the TSG data are:

DATE	Date  (string ‘YYYYMMDDHHMMSS’)

DAYD	Julian day

LATX	Latitude in decimal degree

LONX	Longitude in decimal degree

CNDC	Conductivity measured by the TSG

SSPS	Salinity computed by using CNDC and SSJT

SSJT	Temperature measured by the TSG

SSTP	Temperature measured with a precise sensor (SBE 38 for example)

SPDC	Speed ship (in knots).

FLOW	Flow of pipe that feeds the TSG

PRES	Water pressure in the TSG

Secondary variables based on these main variables are built. Some example for the SSPS variable:

SSPS_QC	Quality code assigned to SSPS

SSPS_CAL	Calibrated SSPS time series

SSPS_ADJUSTED	Adjusted SSPS time series

SSPS_ADJUSTED_QC	Quality code assigned to SSPS

SSPS_ADJUSTED_ERROR	Error assigned to adjusted SSPS time series

Note on the differences between XXXX, XXXX_CAL, XXXX_ADJUSTED where XXXX = SSPS, SSJT, SSTP

XXXX is the raw data (SSJT, SSTP) or is computed (SSPS).

A quality code can be assigned to this data: variable XXXX_QC.

Raw data variable are neither modified nor deleted. If a modification is applied a new variable is created.

Calibrated data: XXXX_CAL.

If a calibration is made the variable XXXX.CAL is created but there is no XXXX_CAL_QC variable. The user must use the quality code of the raw variable: XXXX_QC

Adjusted data: XXXX_ADJUSTED

If a correction is made to the time series the variable XXXX.ADJUSTED is created. This variable only contains corrected measurements. If part of the time series is not corrected those measurements are not stored in the adjusted variable, their value is set to NaN (NoT a Number).

An error estimation is assigned to the adjusted data. It is stored in the XXXX_ADJUSTED_ERROR variable.

The quality code, XXXX_ADJUSTED_QC, assigned to the XXXX_ADJUSTED variable is equal to the quality code assigned to the raw data.

N’aurions nous pas du positionner le code qualité à 5 VALUE_CHANGED pour les mesures corrigées ?

To summarize, when you want work with TSG data you must check the existence of the following variables, in that order:

1 – XXXX_ADJUSTED.

If the variable exists and is not empty, correction has been applied to the raw data. An error is given by the XXXX_ADJUSTED_ERROR variable and the quality code by the XXXX_ADJUSTED_QC variable.

2 – If the XXXX_ADJUSTED variable is empty you must check if  XXXX_CAL exists

In that case a calibration was performed. Use then the XXXX_QC variable to get the quality codes assigned to the measurements.

3 – At least if there is no XXXX_ADJUSTED and XXXX_CAL variables, use the XXXX et XXXX_QC variables.

### Discrete sample

Discrete sample variables have the extension _EXT.

DATE_EXT	Date (string ‘YYYYMMDDHHMMSS’)

DAYD_EXT	Julian day

LATX_EXT	Latitude in decimal degree

LONX_EXT	Longitude in decimal degree

SSPS_EXT	Salinity

SSTP_EXT	Temperature

The SSJT time series can be corrected by using SSTP_EXT data.

Derived variables are:

SSPS_EXT_QC	SSPS Quality codes

SSTP_EXT_QC	SSTP Qualitiy codes

n

SSPS_EXT_TYPE	Instrument used to get the samples (4 digits):

ARGO 	ARGO measurements

CTD 	CTD measurements

UNKN	Unknown

WS (Water Sample)  	Water sample analysed in laboratory

etc.

SSTP_EXT_TYPE 	Instrument used to get the samples (4 digits):

ARGO 	ARGO measurements

CTD 	CTD measurements

UNKN	Unknown

etc.

SSPS_EXT_BOTTLE	number or reference of the water sample bottle (WS)

## Input files

TSGQC can read data files with the following format:

- ARGO	p. 43

- ASCII	p. 44

- ASTROLABE	p. 44

- Merchant ship (Logiciel SODA - SOERE SSS)	p. 45

- NetCDF GOSUD	p. 45

Filename extensions are:

TSG data

ASCII	.tsgqc

ASTROLABE	.ast

NetCDF	.nc

SOERE SSS	.lbv

Sample data

ASCII	.btl	for water sample.

.spl 	for every type of discrete measurements (CTD ...)

ARGO	.arg	ARGO data in LOCEAN format.

## Output files

TSGQC can write data files with the following format:

- NetCDF GOSUD - see CORTSG_format_gosud.doc

netCDF files allows the user to keep all the information about the TSG measurements: meta-data, TSG time series, sample data..

This is the default output format

- ASCII

This format can be used for TSG time series and discrete sample. The format is the samee for input and output files.

A description of the format is available p. 44

## ARGO Format (*.arg)

Nicolas Martin (Ingénieur at LOCEAN) developed a software that collocates ARGO floats position and ship track. The collocation is made for TSG data managed by the Coriolis data center. ARGO measurements are kept if they are found at less than 50 km and 5 days.

The internet address of the collocalisation ARGO-TSG site is:

http://www.locean-ipsl.upmc.fr/~TSG-ARGO/

The site is daily updated.

In the ARGO file there is one line per ARGO collocalisation.

1-6 	Date of the ARGO profile: YYYY-MM-DD HH:MI:SS

7  	Longitude of the ARGO profile

8  	Latitude of the ARGO profile

9  	Profiler number

10  	Cycle number

11  	Pressure of the ARGO measurement (*)

12  	Pressure quality

13  	ARGO salinity

14  	Salinity quality code

15  	ARGO temperature

16  	Temperature quality code

17  	Mean time difference between the ARGO et TSG measurements (**)

18  	Mean distance between the ARGO et TSG measurements

19  	Number of TSG salinity

20  	Average of TSG salinity

21  	Standard deviation of TSG salinity

22  	Number of TSG temperature

23  	Average of TSG temperature

24  	Standard deviation of TSG temperature

(*)	By default it is the pressure the closest to 5 m, between 0 and 10 m and with quality code of ‘1’ or '2'

(**)	Less than R km and at +/-J days (by default R = 50 km and J = 5 days)

Important

TSGQC adds the mean time difference between the ARGO and TSG measurements (column n°17) to the date of the ARGO profile. It is that date which is stored in netCDF and *.spl files.

## ASTROLABE Format (*.ast)

Ce format a été créé pour le TSG du navire des Terres Australes ASTROLABE par Elodie Kestenare (UMR LEGOS)

## ASCII Format (*.btl, *.spl, *.tsgqc)

File with ASCII format have the same structure for TSG (*.tsgqc) and sample files (*.btl, *.spl):

TSG data

%PLATFORM_NAME TOUCAN

%HEADER YEAR MNTH DAYX hh mi ss LATX LONX SSPS SSPS_QC SSJT SSJT_QC SSTP SSTP_QC

2007 07 18 14 20 32   49.9168510   -3.0475171 35.300 0 15.702 0    NaN 9

2007 07 18 14 25 32   49.9076500   -3.0808330 35.290 0 16.028 0    NaN 9

2007 07 18 14 30 32   49.8984489   -3.1142170 35.281 0 15.997 0    NaN 9

2007 07 18 14 35 32   49.8893013   -3.1478500 35.273 0 16.130 0    NaN 9

2007 07 18 14 40 32   49.8801155   -3.1815331 35.260 0 16.170 0    NaN 9

2007 07 18 14 45 32   49.8716011   -3.2158000 35.262 0 16.296 0    NaN 9

There can be one or more header line. They must begin with the ‘%’ character. Metadata can be written in the header lines. In that case you must first indicate the variable name listed in the netCDF format (see CORTSG_format_gosud.doc)

The header line beginning with %HEADER is absolutely necessary. This line describes the variables available in the file.

The 6 first variables are necessary (in that order):

- year	YEAR

- month	MNTH

- day	DAYX

- hour	hh

- minutes	mi

- seconds	ss

Then the order of the netCDF variables is not imposed. In the preceding example the file contains the following variables:

LATX	Latitude

LONX	Longitude

SSPS	TSG salinity

SSPS_QC	Salinity quality code

SSJT	TSG temperature

SSJT_QC 	TSG temperature quality code

SSTP	Temperature sensor

SSTP_QC 	Temperature sensor quality code

In the preceding example the name of the ship where the TSG is installed is assigned to the netCDF variable 'PLATFORM_NAME'.

Avertissement

*.spl files can be used:

- to load discrete samples

- to export in a unique file discrete samples data.

TSGQC adds the mean time difference between the ARGO and TSG measurements (column n°17) to the date of the ARGO profile. It is that date which is stored in netCDF and *.spl files.

## netCDF format (*.nc)

netCDF files allows the user to keep all the information about the TSG measurements: meta-data, TSG time series, sample data..

The format is described in the document CORTSG_format_gosud.doc

## SODA format (*.lbv)

This format is used for the acquisition of TSG data onboard merchant ships managed by the SOERE SSS.

The acquisition software SODA is developed under LabView by David Varillon (IRD - US191 ‘IMAGO’ - Nouméa). The software documentation: SODA_User_manual.pdf is available on demand..

## NUKA ARTICA format (*.transmit)

Ce format a été créé pour le TSG du navire commercial Nuka Artica par ....

## Climatology file format

TSGQC uses the climatology WOA01 and WOA05 (WORLD OCEAN ATLAS 2001 and 2005). The documentation is available at this address:

ftp://ftp.nodc.noaa.gov/pub/WOA01/readme.pdf

TSGQC uses reduced netCDF files containing only the climatology at 0 and 5 m.

Data are available in 3 files: annual, seasonal, monthly

The available variables are :

WOA01_MEAN_SSTP: temperature

WOA01_MEAN_SSPS: salinity

WOA01_MEAN_DOX: oxygen

And the standard deviation (WOA01_STD_SSTP) and the number of observations used to build the climatology (WOA01_OBS_SSTP).

In the netCDF files, the WOA01_TIME dimension will take the following values:

WOA01_TIME = 1 for the annual file

WOA01_TIME = 4 for the seasonal file

WOA01_TIME = 12 for the monthly file

Seasonal file (command ncdump):

$ ncdump -h woa01_seasonal_surf.nc

NetCDF woa01_seasonal_surf {

dimensions:

WOA01_TIME = 4 ;

WOA01_DEPH = 1 ;

WOA01_LATX = 180 ;

WOA01_LONX = 360 ;

variables:

float WOA01_TIME(WOA01_TIME) ;

float WOA01_DEPH(WOA01_DEPH) ;

WOA01_DEPH:long_name = "Depth" ;

WOA01_DEPH:units = "meters" ;

WOA01_DEPH:_FillValue = -100.f ;

float WOA01_LATX(WOA01_LATX) ;

WOA01_LATX:long_name = "Latitude" ;

WOA01_LATX:units = "degree_north" ;

WOA01_LATX:_FillValue = -100.f ;

float WOA01_LONX(WOA01_LONX) ;

WOA01_LONX:long_name = "Longitude" ;

WOA01_LONX:units = "degree_east" ;

WOA01_LONX:_FillValue = -100.f ;

float WOA01_MEAN_SSTP(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_MEAN_SSTP:long_name = "Sea Surface Temperature mean" ;

WOA01_MEAN_SSTP:units = "degre Celcius" ;

WOA01_MEAN_SSTP:_FillValue = -100.f ;

float WOA01_MEAN_SSPS(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_MEAN_SSPS:long_name = "Sea Surface Salinity mean" ;

WOA01_MEAN_SSPS:units = "PSU" ;

WOA01_MEAN_SSPS:_FillValue = -100.f ;

float WOA01_MEAN_DOX1(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_MEAN_DOX1:long_name = "Sea Surface Dissolved oxygen mean" ;

WOA01_MEAN_DOX1:units = "ml/l" ;

WOA01_MEAN_DOX1:_FillValue = -100.f ;

float WOA01_STD_SSTP(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_STD_SSTP:long_name = "Sea Surface Temperature standard deviation" ;

WOA01_STD_SSTP:units = "degre Celcius" ;

WOA01_STD_SSTP:_FillValue = -100.f ;

float WOA01_STD_SSPS(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_STD_SSPS:long_name = "Sea Surface Salinity standard deviation" ;

WOA01_STD_SSPS:units = "PSU" ;

WOA01_STD_SSPS:_FillValue = -100.f ;

float WOA01_STD_DOX1(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_STD_DOX1:long_name = "Sea Surface Dissolved oxygen standard deviation" ;

WOA01_STD_DOX1:units = "ml/l" ;

WOA01_STD_DOX1:_FillValue = -100.f ;

short WOA01_OBS_SSTP(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_OBS_SSTP:long_name = "Sea Surface Temperature number observation" ;

WOA01_OBS_SSTP:units = "none" ;

WOA01_OBS_SSTP:_FillValue = -100s ;

short WOA01_OBS_SSPS(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_OBS_SSPS:long_name = "Sea Surface Salinity number observation" ;

WOA01_OBS_SSPS:units = "none" ;

WOA01_OBS_SSPS:_FillValue = -100s ;

short WOA01_OBS_DOX1(WOA01_TIME, WOA01_DEPH, WOA01_LATX, WOA01_LONX) ;

WOA01_OBS_DOX1:long_name = "Sea Surface Dissolved oxygen number observation" ;

WOA01_OBS_DOX1:units = "none" ;

WOA01_OBS_DOX1:_FillValue = -100s ;

// global attributes:

:description = "WOA01 CLIMATOLOGY Surface Levitus seasonal" ;

:author = "J Grelet IRD US191 Brest" ;

:date = "29 Jan 2008" ;

## Glossaire

GOSUD	Global Ocean Surface Underway Data

www.ifremer.fr/gosud/

SOERE SSS	Systèmes d’Observation et d’Expérimentation pour la Recherche en Environnement Sea Surface Salinity

http://www.legos.obs-mip.fr/fr/observations/sss/

ROSCOP	Report of Observations/Samples Collected by Oceanographic Programmes

http://www.ices.dk/Ocean/roscop/index.asp

Varillon D., Shipboard Oceanographic Data Acquisition - S.O.D.A Version 1.00 - Manuel d’utilisateur, 2009

Varillon D., Shipboard Oceanographic Data Acquisition - S.O.D.A Version 1.00 - Manuel d’utilisateur, 2009

Varillon D., Shipboard Oceanographic Data Acquisition - S.O.D.A Version 1.00 - Manuel d’utilisateur, 2009